Laboratory - General

1. Use appropriate personal protective equipment and procedures for safe laboratory work.

a. Wear gloves when handling samples and LC-MS reagents, materials and components.

b. Label all containers with clear descriptions of contents, dates and preparer initials.

c. Perform all work in accordance with institutional safety procedures.

2. Maintain an organized, neat and clean work space.

a. Keep all reagent bottles, pipet boxes and other containers closed.

b. Clean up after all work.

3. Maintain sufficient quantities of all laboratory supplies.

a. General supplies: Kimwipes, gloves, pipet tips, conical tubes, Eppendorf tubes, etc.

b. LC supplies: PEEK tubing/unions/fittings, columns, fused silica, autosampler vials/septa/caps, 	valve rotor seals, etc.

c. MS supplies: performance standards, ESI spray tips, ion transfer tube, pump oil, etc.

4. Collect and dispose of LC solvent waste in an efficient and timely manner.

a. Before the LC waste collection bottles are full, pour the liquid into the EHS 4-liter plastic liquid 	waste bottles.

b. When the EHS bottles are full, schedule an EHS waste pickup and bottle replacement.

LC-MS Procedures

1. Daily

a. Check the liquid nitrogen Dewar (date, pressure, % capacity).

(1) A spare high-pressure Dewar should be available always.

(2) Label the Dewar tag with the usage start date and note the length of time between 			replacements.

b. Check the MS tune view Instrument Status and LC displays for error messages.

c. Check and record the vacuum pressures on the MS tune view.

d. Check the LC solvent levels, refill as needed, and record date and initials on the bottles.

e. Ensure that the solvent waste drain lines are open and solvent flows when LC pumps are on.

f. Ensure that solvent is flowing through ESI spray tip.

g. Record information on all LC and MS use, maintenance and service in the instrument log 	book.

(1) Instrument use records should include date, user name, MS pressures, instrument performance measures (see below), LC column back pressures, and error messages.

h. If mass calibration is required, see the procedure below.

2. Weekly

a. Check the mechanical pump for oil level and condensation, note in log book, and fill with oil 	and/or ballast as needed (see below).

b. Clean the ion transfer tube if needed (lower fore vacuum pressure).

c. Check the liquid nitrogen Dewar and replace if needed.

LC-MS Operation

1. Prepare for an LC-MS analysis.

LC

a. Purge the plumbing for both pumps and flowmeter, if flow has been off.

b. Prime the syringe and wash the fluidics system.

c. Set the column compartment temperature to 40 C.

d. Start the pumps with appropriate solvent compositions, and check that the column back pressures are appropriate.

MS

a. Ensure that the LC solvent is flowing through the ESI spray tip.

b. Position the spray tip near the MS heated transfer tube to prepare to spray.

c. In the MS tune view, turn the instrument from Standby to On and check the ion signal level 	and stability.  Adjust the tip position and voltage to obtain a stable spray (<~10% TIC variation).

2. Perform an analysis of a standard to assess instrument performance prior to sample analyses.

a. Analyze 2uL of a 5fmol/uL solution of the Pierce Peptide Retention Time Calibration (PRTC) mixture using the PRTC method.

b. Evaluate the data, record measures according to the example below, and ensure appropriate performance.

Peptide			m/z meas.	ppm error	abundance	RT

SSAAPPPPPR*		493.7675	+1.62 ppm	8.5e7		20.5’

SAAGAFGPELSR*	586.7994	+1.53 ppm	1.4e8		25’

NGFILDGFPR*		573.3018	+1.22 ppm	3.1e7		29.5’

3. Perform sample analyses.

a. Record sample information in logbook and note instrument performance and any problems.

*note that for spring-loaded inserts, you must add > 20 µL standard or sample into the insert, put insert into vial and flick vial to remove any residual air bubbles. Air bubbles will result in inconsistent injections and compromise reproducibility of data.

4. When finished with sample analyses, set the system in the Standby condition.

a. In the MS Tune window, click on the “On/Standby” button and put the MS in the Standby mode.

b. Set the LC conditions appropriately.

c. The “NC” pump should remain on with an ~50-100nL/min flow rate at 50:50 A:B through the analytical column.

d. The loading pump flow rate should remain on with an ~0.5-1uL/min flow rate at 50:50 A:B through the sample loop and trap column.

e. Leave the power on for all modules.

Maintenance

LC (see Thermo Scientific Dionex UltiMate 3000 RSLCnano manuals for more information.)

1. Prepare new LC solvents using LC-MS grade solvents and chemicals.

a. Prepare A as 0.1% formic acid/100% water, B as 80% acetonitrile/20% water/0.1% formic acid, 	and the loading pump (LP) solvent as 99% water/1% acetonitrile/0.1% formic acid.

b. Prepare the seal wash as 10% methanol/90% water, and the autosampler wash as 50% 	isopropanol/50% water/0.1% formic acid.

c. Purge A, B, and LP solvents.

d. Perform Pressure Transducer Test.

e. Perform Viscosity Measurement.

MS (citations to pages in the Exactive Series Operating Manual, Rev A, Sep 2014)

1. Use the following Thermo schedule as a guide for recommended maintenance tasks.

2. Ion transfer tube (capillary) cleaning (p. 8-23)

a. This component can become blocked, which will manifest in a lower than usual Fore Vacuum 	pressure reading.

b. Turn off the LC flow and place the MS electronics switch in the Service Mode.  Wait ~60min. 	for the tube to cool.

c. If using the HESI source, remove the ion source housing and ion sweep cone.

d. Remove the tube using the removal tool by turning it counterclockwise until loose and then 	pulling out slowly.

e. Sonicate the tube in a 50:50 methanol:water solution in a small graduated cylinder.

f. Remove residual solvent with dry nitrogen (liquid nitrogen Dewar vent).

g. Insert the tube into the MS using the tool turning clockwise and slowly pushing it in.

h. Replace other source components, switch to Operating Mode, heat the capillary and check 	the pressure.

*for the lipid QE+, first sonicate with 50:50 IPA/water solution for 15 min followed by sonication for 15 min in 50:50 methanol:water as described above. Clean both sweep plate and ion transfer capillary.

3. Mechanical pump maintenance (pp. 6-20, 8-9)

a. Purge the pump oil (ballast the pump) once per month

(1) Turn the ballast valve 90 degrees so that the holes are aligned.

(2) Ballast for ~30 minutes.

(3) Turn the valve 90 degrees to close off the holes.

b. Change the pump oil

(1) Vent the system, allow pump to cool, and disconnect plumbing if necessary.

(2) Drain the pump oil into a pan and pour the used oil into a container for disposal.

(3) Add new oil, reconnect the vacuum tubing if necessary, and restart the system.

(4) Change lipid QE+ pump oil every 3 months.

(5) Change proteomics QE+ pump oil every 6 months.

4. Use the following Thermo schedule as a guide for recommended calibration tasks.

5. MS calibration

a. Install the HESI source.

b. Fill syringe with Thermo LTQ Velos Positive Ion Calibration Mixture and flush the tubing.

c. Start the syringe pump at a 1-5 uL/minute flow rate.

d. Load the HESI source tune file and turn the MS “On”.

e. Adjust flow rate and spray voltage to obtain a stable spray (<10% TIC variation).

f. Start the Positive Ion Calibration, which is located on the Calibrate tab on the left side of the 	MS tune view.

g. Perform other calibrations and evaluations as needed.

Computer

1. Save MS method and data files in appropriate, user- or project-named folders.

2. Use file names that include the date, user initials, and sample and method information.

3. Copy data files to other media or computers for data analysis and storage as archival records.

4. Periodically check the amount of free disk space and remove files as necessary.

Consumable materials and replacement parts

1. Ion transfer tube, Thermo, P/N 70005-20606

2. Mechanical pump oil,

3. Thermo LTQ Velos positive ion calibration mixture

4. Thermo Pierce Peptide Retention Time Calibration Mixture, 0.5pmol/uL , 50uL, P/N 88320

Dilute to 5fmol/uL in LC solvent A

5. Stainless steel nanobore emitters, 40mm, OD 1/32”, P/N ES542

6. Water, acetonitrile, methanol; 4-liter bottles, Optima grade

7. Formic acid, ThermoFisher Pierce LC-MS grade, 50mL, P/N 85178

8. Nanotrap column, Acclaim PepMap100 C18, 5um, 100A, 100umx2cm, nanoViper, P/N 164564

9. Analytical column, Acclaim PepMap RSLC C18, 2um, 100A, 75umx150mm, nanoViper, P/N 164769

10. Autosampler vials

11. Autosampler vial inserts

12. Autosampler vial caps, ChemGlass Life Sciences, Natural PTFE/Off White silicone low bleed, blue 9mm screw cap, P/N CV-3842-CERT

13. LC valve seals - ?